DIAMONDFLOW™ TECHNICAL DATA  SHEET

| Clinical Indications | Universal Esthetic Restorative Material: Conservative Class I, Pedodontic Fillings, Class V Restorations, Cervical Erosion Repairs, Liner, Base, Tunnel Preps., & Block-out Material |
| --- | --- |
| Composition | PolyCrystalline NanoCeram |
| Mode of Cure | Visible Light Cure (VLC) |
| Delivery System | 5g. Luer-Lock (L.L.) Syringe & 15 gauge 45 º < bent applicator L.L. Needles |

| #### PROPERTY ------------------- STD. SPEC. # | #### DIAMONDFLOW | #### ISO/ADA….. #### Criteria | Tooth Enamel/Dentin | ##### Gold | ##### Porcelain |
| --- | --- | --- | --- | --- | --- |
| Diametral Tensile Strength (MPa) 4049(I) / 27(A) | 56 | 34 | 70 | N/A | 35 |
| Biaxial Flexural Strength (MPa) 4049(I) / F394-78(T) | 176 | 52 | 145 | 186 | 68 |
| Compressive Strength (MPa) 4049(I) / 27(A) | 375 | N/A | 393/290 | 166 | 379 |
| Fracture Toughness (Kicsr) (MPa m-1/2) 1304(T) | 2.0 | N/A | 1.7 | 2.2 | 0.5 |
| Wear Resistance (µm/year) 4049(I)/(C)/ (D) | 10 | 50 | 3-5 | 3-5 | “ wears opposing dentition “ |
| Water Sorption (%wt.) (µg/cm3) 4049(I) / 27(A) | 0.50 | 0.70 | N/A | N/A | N/A |
| Polymerization Shrinkage(% linear) 4049(I) | 0.65 | N/A | ------ | 1.56 | 17 |
| Linear Coefficient Thermal Expansion LCTE - (ppm) (mm/mm/0C x 10-6) (D) | 24 | N/A | 11.4 | 14.2 | 14 |
| Ra-Profilometry Average Surface Roughness (nm)(D) | 1 | N/A | 5 | 15 | 15 |
| VHN Surface Hardness @ 3mm 4049(I) / 27(A) | 99 | 60 | 210/80 | 80 | 475 |
| Depth of Cure (mm) 4049(I) / 27(A) | 5.2 | 2.0 | N/A | N/A | N/A |
| Biocompatability 10993(I) / 41(A) | Excellent | Moderate | ------ | Excellent | Moderate |
| Cytotoxicity (%) 10993(I) / 41(A) | 0 | 40 | ------ | 0 | 25 |
| Polishability (D) | Very Good | N/A | Very Good | Very Good | “ N/A-Glazed “ |
| Shades (Vita) (V) | 5 Vita | N/A | ------ | 1 | Vita |
| Particle Size Microns (μm), Nano (nm) (D) | 0.25 μm (+) 7 nm | N/A | ------ | N/A | N/A |
| Filler Weight (%), Fw (D) | 70.0 | 50.0 | ------ | ----- | ----- |
| Filler Volume (%), Vf (D) | 61.0 | 30.0 | ------ | ----- | ----- |
| Resin / Matrix (D) [Radiopaque, (X-Ray detectable) & Fluorescent] | Crystalline - PEX | N/A | Crystalline - HA | ----- | ----- |
| DiamondFlow™ Features & Benefits STANDARDS – SPECIFICATION # ISO-4049 ~ 4049(I) ISO-10993 ~ 10993(I) ADA-27 ~ 27(A) ADA 41~ 41(A) ASTM 1304 ~ 1304(T) ASTM-ANSI F394-78 ~ F394-78(T) VITA SHADE ~ (V) CDA ~ (C) DRM ~ (D) | - Cyclic Fatigue Resistant, High Fracture Toughness, Shock Absorbing - Exceptional Fracture Resistance: Tensile, Shear, Compressive & Flexural. - Wear Resistance Similar to Natural Tooth - Ultimate Reinforcement and Handling Characteristics, Unique - Controlled Flowable-Technology - Thixotropic(Shape-Memory Recall, Anti-Slumping Properties), Easy to Model, Carve, Shape and Anatomically Contour - Superb Color Stability and Accurate Shade Match - Easy Access, infiltrates the Most Difficult Preps., Enhanced Contact Angle Wettability / Adaptation to Accute Transitional Prep. Line Angles - Superb Bonding Integration for Maximum Adhesion - (5) Popular Colors: Vita Equiv. - A1, A2, A3, A3.5& A2.5(U) - Delivery Systems, available in 5g. Luer-Lock (L.L.) Syringe & 15 gauge 45 º < bent applicator L.L. Needles - Refractive-Index, Chameleon-Like Effect - Specular Reflectance, Gloss & Luster Value - Ultimate Longevity; Negligible Sorption, Absence of Solubility - Excellent Marginal Integrity & Optimum Shock Absorption Characteristics @ the Liner/Base-Filling Interface (Micro-Elastic), Low Shrinkage - Biocompatible & Non-Cytotoxic |  |  |  |  |