DIAMONDLINK™ TECHNICAL DATA  SHEET

| Clinical Indications | Universal Esthetic Bonding/Luting Cement Fast-Cure Material - Enamel, Dentin, Resin, Ceramic & Metal Substrates - For Metal-Free Crown & Bridge, Inlays/Onlays, Laminate-Veneers, Implant-Superstructures, Pins, Post/Core, Fiber Posts, Splints, Ortho-Brackets, Pit & Fissure Sealant, Flowable Cavity Liner/Base & Temporary Crown Material |
| --- | --- |
| Composition | Poly-Crystalline Nano-Ceram |
| Mode of Cure | Dual-Cure : Visible Light Cure (Base) + Self/Auto-Cure (Catalyst), Cure - Light, Auto or Dual |
| Delivery System | 5g. Luer-Lock (L.L.) Syringe & 18-22 gauge applicator L.L. Needles |

| #### PROPERTY ------------------- STD. SPEC. # | #### DIAMOND #### LINK | #### ISO/ADA….. #### Criteria | Tooth Enamel/Dentin | ##### Gold | ##### Porcelain |
| --- | --- | --- | --- | --- | --- |
| Diametral Tensile Strength (MPa) 4049(I) / 27(A) | 59 | 34 | 70 | N/A | 35 |
| Biaxial Flexural Strength (MPa) 4049(I) / F394-78(T) | 145 | 52 | 145 | 186 | 68 |
| Compressive Strength (MPa) 4049(I) / 27(A) | 350 | N/A | 393/290 | 166 | 379 |
| Fracture Toughness (Kicsr) (MPa m-1/2) 1304(T) | 1.6 | N/A | 1.7 | 2.2 | 0.5 |
| Wear Resistance (µm/year) 4049(I)/I/ (D) | 7 | 50 | 3-5 | 3-5 | “ wears opposing dentition “ |
| Water Sorption (%wt.) (µg/cm3) 4049(I) / 27(A) | 0.45 | 0.70 | N/A | N/A | N/A |
| Polymerization Shrinkage(% linear) 4049(I) | 0.65 | N/A | ------ | 1.56 | 17 |
| Linear Coefficient Thermal Expansion LCTE – (ppm) (mm/mm/0C x 10-6) (D) | 22 | N/A | 11.4 | 14.2 | 14 |
| Ra-Profilometry Average Surface Roughness (nm)(D) | 2 | N/A | 5 | 15 | 15 |
| VHN Surface Hardness @ 3mm 4049(I) / 27(A) | 97 | 60 | 210/80 | 80 | 475 |
| Depth of Cure (mm) 4049(I) / 27(A) | 5.2 | 2.0 | N/A | N/A | N/A |
| Film Thickness, Microns (μm) 3107(I)/ 8(A)/ 9(A) | 13 | 40 | N/A | N/A | N/A |
| Cure Kinetics, Tw, Tg, Ts [ Work(w), Gel(g), Set(s) - Time] - Min’ Sec’’ @ 370C 3107(I)/ 8(A)/ 9(A) Self/Auto Cure Mode (Cat. +Base) (1:1 Ratio) | Tw - 1’30’’ Tg - 3’30’’ Ts - 5’30’’ | 1’00’’minima- 5’00’’maxima- 9’00’’maxima | N/A | N/A | N/A |
| Bond-Strength (MPa) 3107(I)/ 8(A)/ 9(A) (*) DLINK2-F/C-DBOND ---( DCROWN / Tooth ) | * 27 | 4 | 17/15 | 6/8/10 - ** ZP/PC/GI | 4/6/8 - ** ZP/PC/GI |
| Biocompatability 10993(I) / 41(A) | Excellent | Moderate | ------ | Excellent | Moderate |
| Cytotoxicity (%) 10993(I) / 41(A) | 0 | 40 | ------ | 0 | 25 |
| Polishability (D) | Very Good | N/A | Very Good | Very Good | “ N/A-Glazed “ |
| Shades (Vita) (V) | 3 Vita | N/A | ------ | 1 | Vita |
| Particle Size Microns (μm), Nano (nm) (D) | 0.25-0.40μm - (+) 7 nm | N/A | ------ | N/A | N/A |
| Filler Weight (%), Fw (D) | 74.0 | 50.0 | ------ | ----- | ----- |
| Filler Volume (%), Vf (D) | 64.0 | 30.0 | ------ | ----- | ----- |
| Resin / Matrix (D) [Radiopaque, (X-Ray detectable) & Fluorescent] | Crystalline -PEX | N/A | Crystalline - HA | ----- | ----- |
| DiamondLink™ Features & Benefits STANDARDS – SPECIFICATION # ISO-4049 ~ 4049(I) ISO-10993 ~ 10993(I) ISO-3107 ~ 3107(I) ADA-27 ~ 27(A) ADA-8 ~ 8(A) ADA-9 ~ 9(A) ADA 41~ 41(A) ASTM 1304 ~ 1304(T) ASTM-ANSI F394-78 ~ F394-78(T) VITA SHADE ~ (V) CDA ~ (C) DRM ~ (D) ** ZP(ZincPhosphate), PC(Polycarboxylate, GI(Glass Ionomer) Cements – Bonded to Gold/Porcelain Crown * DiamondCrown Cemented w/ DiamondLink Cement & DiamondBond Adhesive | - Cyclic Fatigue Resistant, High Fracture Toughness, Shock Absorbing - Exceptional Fracture Resistance: Tensile, Shear, Compressive & Flexural. - Wear Resistance Similar to Natural Tooth, Ideal as Cement, Liner/Base & Sealant - Ultimate Reinforcement and Handling Characteristics, Controlled Flow-Technology - Ultra Low Film Thickness ; Tough & Thin Interface for Optimum Crown/Abutment Tooth Fit Integrity and Marginal Seal Impermeability - Superb Color Stability, Stain Resistance and Accurate Esthetic Shade Match - Easy Access, infiltrates the Most Difficult Preps., Enhanced Contact Angle Wetability / Adaptation to Acute Transitional Prep. Line Angles - Superb Bonding Integration for Maximum Adhesion - (3) Popular Colors: Vita Equiv. - A2 (Universal/Natural), B1(Translucent), B3(Yellow) - Delivery Systems, avail. in 5g. Luer-Lock (L.L.) Syringe & 18-22 gauge 45 º L.L. Needles - Refractive-Index, Chameleon-Like Effect - Specular Reflectance, Gloss & Luster Value - Ultimate Longevity; Negligible Sorption, Absence of Solubility - Excellent Marginal Integrity & Optimum Shock Absorption Characteristics @ the Liner/Base-Filling &/or Crown-Abutment Tooth Interface&(Micro-Elastic), Low Shrinkage - Biocompatible & Non-Cytotoxic - Strong, Tough, Esthetic & Tissue Compatible Temporary Crown Material |  |  |  |  |